## device description
Test report on MightyWatt electronic load for Arduino
2015-12-03
Jakub Polonský (kaktus circuits)
MightyWatt
standard 30V/10A/70W unit
revision 2.5
manufactured 2015-11
standard terminals
firmware 2.5.1
Power supply
PSD30/3C linear laboratory power supply
30V/3A
oscilloscope noise measurements: 2× Panasonic Eneloop Pro 1.2 V NiMH batteries in series
Ammeter
Multimeter Keysight 34461A
auto range, max 3A, 10 PLC integration
Oscilloscope
Rigol DS1052E
probe 1:1
Ambient conditions
21 °C

## Arduino Uno
Calibration
DAC | ADC | Actual I, A |  | IDAC_SLOPE | 10492.1088
40 | 28 | 0.1008 |  | IDAC_INTERCEPT | -1.558
320 | 308 | 0.81816
600 | 588 | 1.5355 |  | IADC_SLOPE | 10477.1187
880 | 868 | 2.2527 |  | IADC_INTERCEPT | 30.3093
1160 | 1150 | 2.9697
Physical ADC LSB | 10.2316 | mA
Oversampled ADC LSB | 2.5579 | mA
Physical DAC LSB | 2.5615 | mA
Measurement
Set I, A | MightyWatt measured I, A | External ammeter I, A | Relative deviation of set current | Relative deviation of measured current | Relative deviation of set current relative to full range (10A) | Relative deviation of measured current relative to full range (10A) | Deviation of set current, LSB | Deviation of measured current, physical LSB | Deviation of measured current, oversampled LSB
0.1 | 0.101 | 0.0982 | 1.833% | 2.851% | 0.01716% | 0.02672% | 0.7027 | 0.27366 | 1.0947
0.25 | 0.255 | 0.2494 | 0.2406% | 2.245% | 0.005719% | 0.05345% | 0.23423 | 0.54733 | 2.1893
0.5 | 0.5 | 0.498 | 0.4016% | 0.4016% | 0.01906% | 0.01909% | 0.78078 | 0.19547 | 0.78189
1 | 1.001 | 1.0001 | -0.009999% | 0.08999% | -0.0009531% | 0.00859% | -0.039039 | 0.087963 | 0.35185
1.5 | 1.498 | 1.4997 | 0.02% | -0.1134% | 0.002859% | -0.01623% | 0.11712 | -0.16615 | -0.66461
2 | 1.994 | 1.9992 | 0.04002% | -0.2601% | 0.007625% | -0.04963% | 0.31231 | -0.50823 | -2.0329
2.5 | 2.495 | 2.4985 | 0.06004% | -0.1401% | 0.0143% | -0.03341% | 0.58558 | -0.34208 | -1.3683
3 | 2.997 | 2.9975 | 0.0834% | -0.01668% | 0.02383% | -0.004772% | 0.97597 | -0.048868 | -0.19547
3 | 2.997 | 3.008 | -0.266% | -0.3657% | -0.07625% | -0.105% | -3.1231 | -1.0751 | -4.3004
21 | °C | Temperature coefficient
88 | °C | 52.2824 | ppm / °C
|  | 50.0417 | ppm / W

## Arduino Due
Calibration
DAC | ADC | Actual I, A |  | IDAC_SLOPE | 10494.3909
40 | 15 | 0.10085 |  | IDAC_INTERCEPT | -1.6263
320 | 296 | 0.8182
600 | 577 | 1.5357 |  | IADC_SLOPE | 10445.8881
880 | 859 | 2.2531 |  | IADC_INTERCEPT | 63.1146
1160 | 1140 | 2.9704
Physical ADC LSB | 2.5503 | mA
Physical DAC LSB | 2.5621 | mA
Measurement
Set I, A | MightyWatt measured I, A | External ammeter I, A | Relative deviation of set current | Relative deviation of measured current | Relative deviation of set current relative to full range (10A) | Relative deviation of measured current relative to full range (10A) | Deviation of set current, LSB | Deviation of measured current, LSB
0.1 | 0.098 | 0.09818 | 1.854% | -0.1833% | 0.01734% | -0.001723% | 0.71035 | -0.070581
0.25 | 0.249 | 0.2494 | 0.2406% | -0.1604% | 0.005717% | -0.003829% | 0.23418 | -0.15685
0.5 | 0.499 | 0.4979 | 0.4218% | 0.2209% | 0.02001% | 0.01053% | 0.81964 | 0.43133
1 | 1.001 | 1.0001 | -0.006% | 0.09399% | -0.0005717% | 0.008999% | -0.023418 | 0.36859
1.5 | 1.501 | 1.4997 | 0.02% | 0.08668% | 0.002859% | 0.01245% | 0.11709 | 0.50975
2 | 2.001 | 1.9992 | 0.04002% | 0.09004% | 0.007623% | 0.01723% | 0.31224 | 0.70581
2.5 | 2.498 | 2.4986 | 0.05603% | -0.02401% | 0.01334% | -0.005744% | 0.54643 | -0.23527
3 | 2.998 | 2.9974 | 0.08674% | 0.02002% | 0.02478% | 0.005744% | 1.0148 | 0.23527

## Noise
Current noise
1 Ohm 0.1% resistor
1 A
Oscilloscope | Multimeter
| 10 PLC | 1 A range
| Average current | 1.0004 | A
| Standard deviation | 1.32e-05 | A
| Relative standard deviation | 0.00132%
| Minimum value | 1.0003 | A
| Maximum value | 1.0004 | A
| Number of samples | 600
> 3 Hz
| DC I, A | AC I, mA (RMS) | AC/DC ratio
| 3 | 1.67 | 0.05567%
| 1 | 0.63 | 0.063%
| 0.5 | 0.14 | 0.028%
| 0.1 | 0.235 | 0.235%